## Cover
Tidepool Loop 1.0.0 Requirements, Traceability, Hazard Analysis and Verification Reports
Automatically generated on {timestamp}
Corresponds with Tidepool Loop 1.0.0 (build {build_number})
This document is intended to show the detail of the Tidepool’s Quality Management System for Tidepool Loop. The document was automatically generated from Tidepool’s quality management system. The source documents are in Jira, our project management system. Information was pulled from these projects: • Tidepool Loop Functional Requirements - TLFR • Tidepool Loop Risks - TLR • Tidepool Loop Work Tickets - LOOP Tabs in this spreadsheet: • Software Requirements: This spreadsheet tab contains all software requirements for Tidepool Loop, as derived from the Tidepool Loop Functional Requirements (TLFR) Jira project. • Traceability Summary: This spreadsheet tab summarizes traceability from requirements to implementation and verification summary. It does not include risk analysis for each requirement. • Traceability Report: This spreadsheet tab demonstrates traceability to and from all Tidepool Loop functional requirements. It has columns for each requirement, associated risks, design and development, verification tests, and test status. • For more information on Tidepool Loop’s requirements, please see Section 16 of the main pre-marketing notification document. • Full Traceability Report: This spreadsheet tab includes all of the above information, but also includes traceability to all risk management issues. • Hazard Analysis: This spreadsheet tab demonstrates Tidepool’s Risk and Hazard analysis of Tidepool Loop. Columns are included for each risk, including a summary, the potential harm and hazard category, risk assessment prior to mitigation, details of any mitigation, and risk assessment post-mitigation. • Unresolved Defects: This spreadsheet lists all unresolved defects and anomalies (open bugs) for Tidepool Loop 1.0.0. • Automated Tests: This spreadsheet lists additional verification testing performed via test automation methods.

## Glossary
Requirements Glossary
| Term | Definition
A | ACE_PUMP / pump | “Alternate Controller Enabled” insulin pump.
| ACTIVE_CARBS | The amount of user-entered carbohydrates that have been consumed, but not fully absorbed.
| ACTIVE_INSULIN | The product of the amount of insulin delivered and the percent of INSULIN_ACTIVITY_DURATION remaining.
B | BOLUS/bolus | A one-time dose of insulin given to bring down high glucose or cover an expected rise in glucose from eating.
| BG/Glucose | Glucose measurement.
C | CALIBRATION_GLUCOSE_VALUE | Glucose value measured thru a blood sample to optimize system performance.
| CARBOHYDRATE_TO_INSULIN_RATIO | The amount of insulin needed to cover a given amount of carbohydrates. Expressed in (gU).
| CORRECTION_RANGE | User setting of glucose range the Tidepool Loop algorithm uses to determine corrective actions. The range will have an upper bound and lower bound.
| CORRECTION_TARGET | From t0 to td/2, the correction target is equal to the SUSPEND_THRESHOLD. From td/2 to td, the correction target is a linear blend over time of the SUSPEND_THRESHOLD and the average of the correction maximum and correction minimum.
| CRITICAL_ALERTS | A specific iOS feature that bypasses silencing the phone.
| CRITICAL_ALERT_ENTITLEMENT | iOS implementation of CRITICAL_ALERTS.
| CURRENT_GLUCOSE | The last CGM data point or manual glucose reading that is less than 15 minutes old.
D | DELIVERY_LIMITS | Safety guardrails for your insulin delivery.
| DOSING_SAFETY_THRESHOLD | Equivalent to the user's SUSPEND_THRESHOLD setting for the first half of the INSULIN_ACTIVITY_DURATION, and then linearly increases until it reaches the midpoint of the CORRECTION_RANGE at the end of the INSULIN_ACTIVITY_DURATION.
| DOSING_UNCERTAINTY | Tidepool Loop issues a command to the ACE Pump, such as deliver a BOLUS or temp basal, but does not receive confirmation that the command was either successfully executed or failed to execute.
E | EVENTUAL_GLUCOSE | The predicted glucose value at time t, which is the furthest point in the future that the algorithm can predict and is dependent on the INSULIN_ACTIVITY_DURATION.
F
G | GLUCOSE_FORECAST | Predicted glucose which is an output of the Tidepool Loop Algorithm.
| GLUCOSE_SAFETY_LIMIT | User selected glucose value where Tidepool Loop will suspend the delivery of basal insulin.
H
I | ICGM | Integrated Continuous Glucose Monitor
| ICGM_PILL | Touch sensitive area on Tidepool Loop Home Screen displaying information about connected CGM.
| IN-APP_MODAL | Messages pushed to the user in the form of Notifications, Alerts or Alarms.
| INSULIN_ACTIVITY_DURATION | The amount of time insulin is active in the body. Tidepool Loop sets this time to 6 hours.
| INSULIN_ACTION_PEAK | Time during which insulin is at its maximum effectiveness.
| INSULIN_SENSITIVITY_FACTOR | The expected drop in glucose given one unit of insulin.
| IOB | See ACTIVE_INSULIN
J
K
L | LOOP_NOT_LOOPING | Automated Insulin delivery is not active.
| LOOP_STATUS_PILL | Tappable status icons displayed at the top of the Tidepool Loop app to check in with how the app is working.
M | MAXIUMUM_BASAL_RATE | The maximum amount of basal insulin that the pump can be programmed to deliver by either the user or Tidepool Loop’s algorithm to deliver per hour.
| MINIMUM_BOLUS_INCREMENT | The smallest amount of BOLUS insulin that the pump can be programmed to deliver.
N | NO_GLUCOSE_MSG | Message displayed to the user when there is not a CURRENT_GLUCOSE.
O
P | PREDICTED_GLUCOSE | Forecast of future glucose values that spans from the current time to to td, some time in the future.
| PREDICTIVE_ALERTS | User configured USER_NOTIFICATION based on PREDICATED_GLUCOSE.
| PRE_MEAL_RANGE | You may choose to set a lower temporary glucose Correction Range up to one hour before eating so that you begin a meal in a lower range. This feature can be used before bolusing for meals in an effort to impact post-meal glucose levels. This pre-meal glucose Correction Range will remain active for up to 1 hour after it’s activated, or until a carbohydrate entry has been entered into Tidepool Loop.
| PRE_MEAL_TEMP_ADJUST | Feature telling Tidepool Loop to adjust your basal insulin and lower your glucose CORRECTION_RANGE in advance of your meal. The temp adjust will be in effect for up to one hour or until you cancel it or enter your next Carb Entry.
PUMP_TIMEZONE_OFFSET
Q
R
S | SCHEDULED_BASAL_RATE | The amount of insulin you and your healthcare provider have set your insulin pump to deliver for a specific time segment throughout the day.
| STATUS_INDICATOR | Icon that allows you to see if automation is on and working.
| SUSPEND_THRESHOLD / GLUCOSE_SAFETY_THRESHOLD | User selected glucose value where Tidepool Loop will suspend the delivery of basal insulin.
T | TEMPORARY_BASAL_RATE | The amount of insulin needed beyond the user’s SCHEDULED_BASAL_RATE to keep the PREDICTED_GLUCOSE within the CORRECTION_RANGE.
| TEMPORARY_CORRECTION_RANGE | CORRECTION_RANGE used when the user has enabled PRE_MEAL_TEMP_ADJUST or WORKOUT_TEMP_ADJUST.
| TEMPORARY_STATUS_BANNER/TEMPORARY_STATUS_BAR | Banner notification displayed in the Tidepool Loop app below the status icons to alert you to a temporary status.
| TIDEPOOL_LOOP_PRESCRIPTION | Authorization to begin using Tidepool which will also contain initial Tidepool Loop settings.
| TIDEPOOL_LOOP_PRESCRIPTION_CODE | Unique code from your healthcare provider needed to begin using Tidepool Loop.
U | USER_NOTIFICATION/notification | User-facing notifications communicating important information to Tidepool Loop Users. Notifications can be delivered as Alerts or Alarms.
V
W | WORKOUT_TEMP_ADJUST | Feature telling Tidepool Loop to adjust both your basal insulin and your glucose CORRECTION_RANGE to help you meet your glucose goals during that activity. Workout temp adjust will be in effect for the time you indicate when you activate it or until you cancel it.
X
Y
Z

## Software Requirements
Software Requirements
Req ID | Req Key | Req Feature | Req Description
<<<insert: requirements_list()>>>

## Traceability Summary
Traceability Summary
Req ID | Req Key | Req Feature | Req Description | Story Key | Story Summary | Status
<<<insert: traceability_summary()>>>

## Traceability Report
Traceability Report
Req ID | Req Key | Req Feature | Req Description | Story Key | Story Summary | Test Key | Test Summary | Test Status
<<<insert: traceability_report()>>>

## Full Traceability Report
Full Traceability Report
Req ID | Req Key | Req Feature | Req Description | Story Key | Story Summary | Test Key | Test Summary | Test Status | Risk Key | Hazardous Situation
<<<insert: traceability_report('full': True)>>>

## Risk Scoring
Severity Scoring
Severity Score | Criteria for TRSET outputs (where applicable) | Category | Description
5 |  | Catastrophic | The hazard results in death or life-threatening injury
4 |  | Critical | The hazard results in permanent impairment
| DKAI ≥ 21 hrs |  | The person with diabetes likely has Beta-Hydroxybutyrate (BHB) present in their blood at a level (>4.4 mmol/L) that indicates that they are in Diabetic Ketoacidosis (DKA), though the severity of the DKA could vary between mild, moderate, and severe.
| LBGI > 10 |  | The person with diabetes experiences a hypoglycemic event resulting in death, permanent impairment, or life-threatening injury.
3 |  | Serious | The hazard results in injury or impairment requiring medical intervention, including administration of IV fluids
| 14 ≤ DKAI < 21 hrs |  | The person with diabetes likely has BHB present in their blood at a level (>3.0 mmol/L) that indicates they are likely, though not certainly in DKA.
| 5 < LBGI ≤ 10 |  | The person with diabetes experiences a hypoglycemic event that results in injury or impairment requiring medical intervention, including administration of IV fluids.
2 |  | Minor | The hazard results in temporary injury or impairment not requiring medical intervention.
| 8 ≤ DKAI < 14 hrs |  | The person with diabetes likely has BHB present in their blood at a level (BHB > 1.8 mmol/L) that requires monitoring, but is only a minor risk of DKA.
| 2.5 < LBGI ≤ 5 |  | The person with diabetes experiences a hypoglycemic event that results in temporary injury or impairment not requiring medical intervention.
1 |  | Negligible | The hazard results in inconvenience or temporary discomfort.
| 2 ≤ DKAI < 8 hrs |  | The person with diabetes likely has BHB present in their blood at a level (<1.8 mmol/L) that is not likely to be DKA.
| LBGI ≤ 2.5 |  | The person with diabetes experiences a hypoglycemic event that results in inconvenience or temporary discomfort.
Probability Scoring
Rating | Term | Probability (P) of occurrence of the harm during the duration of insulin action | Description for clarity
5 | Frequent | P ≥ 0.1 | Likely to occur 1 in 10 times or more often.
4 | Probable | .01 ≤ P < 0.1 | Likely to occur between (1 in 10) and (1 in 100) times.
3 | Occasional | .0001 ≤ P < 0.01 | Likely to occur between (1 in 100) and (1 in 10,000) times.
2 | Remote | .000001 ≤ P < 0.0001 | Likely to occur between (1 in 10,000) and (1 in 1,000,000) times.
1 | Improbable | P ≤ 0.000001 | Likely to occur less frequently than 1 in a million times.

## Risk Matrix
Risk Acceptability Matrix
Severity
|  | Negligible (1) | Minor (2) | Serious (3) | Critical (4) | Catastrophic (5)
Probability | Frequent (5)
Probable (4)
Occasional (3)
Remote (2)
Improbable (1)

## Hazard Analysis
Hazard Analysis | [8 empty] | 95 (47.98%) | [5 empty] | 110 (55.56%)
[9 empty] | 80 (40.4%) | [5 empty] | 40 (20.2%)
[9 empty] | 21 (10.61%) | [5 empty] | 0 (0.0%)
[9 empty] | 2 (1.01%) | [5 empty] | 48 (24.24%)
[9 empty] | 198 (100%) | [5 empty] | 198 (100%)
Risk Key | Hazardous Situation | Source | Sequence of Events | Harm | Hazard | Hazard Category | Initial Severity | Initial Probability | Initial Risk Score | Mitigation Key | Mitigation Summary | Mitigation Description | Residual Severity | Residual Probability | Residual Risk Score | Benefit outweigh Risks? Y/N
<<<insert: hazard_analysis()>>>

## Insulin Fidelity
Insulin Fidelity |  |  |  |  | 95 (47.98%) | 110 (55.56%)
[5 empty] | 80 (40.4%) | 40 (20.2%)
[5 empty] | 21 (10.61%) | 0 (0.0%)
[5 empty] | 2 (1.01%) | 48 (24.24%)
[5 empty] | 198 (100%) | 198 (100%)
Risk Key | Hazardous Situation | Hazard | Hazard Category | Initial Severity | Initial Risk Score | Residual Risk Score | Req ID | Mitigation Description | Work Tickets | Verification Tickets
<<<insert: insulin_fidelity('filter': [ 'TLR-548','TLR-552','TLR-576','TLR-577','TLR-578','TLR-579','TLR-581' ])>>>

## Unresolved Defects
Unresolved Defects
Defect Key | Defect Summary | Status | Fix Version
<<<insert: bugs_list('empty': 'No unresolved defects')>>>

## Automated Tests
Automated Tests
Test Report | Test Suite | Test Case | Time | Status
<<<insert: tests_list()>>>